EIDC Supporting information

1-minute NH3 measurements and meteorological parameters from six miniaturized sensors and a Picarro NH3 analyzer at Whim Bog (July to August 2024)

- File 1: 1-minute NH3 measurements and meteorological parameters from six miniaturized sensors and a Picarro NH3 analyzer at Whim Bog (July to August 2024)

This file contains 1-min resolution data for Whim Bog NH3 release system activation, wind speed and direction as well as NH3 concentrations, temperature and RH (not all sensors have them) from six miniaturized NH3 sensors and a Picarro G2103 NH3 analyzer, measured between July 2024 and August 2024.

# Experimental design/sampling regime

This dataset comes from the  first outdoor intercomparison of NH3 sensors in the world, using the Picarro G2103 NH3 analyzer as a reference instrument to compare against the miniaturized NH3 sensors.

The Picarro G2013 NH3 analyser and sensors were installed 8 m away from the NH3 release system at Whim Bog in close proximity to the release point to ensure high concentrations. The experiment was carried out from 24/07/2024 to 28/08/2024. Both the Picarro analyser and the sensors were placed 0.5 m above the boardwalk.

The Picarro analyser was placed inside a weatherproof air-conditioned box with temperature control. The sensor housing (SH) was installed next to the Picarro box. The SH was built for this experiment and consisted of a polycarbonate weatherproof box (24.5 x 20.5 x 9.5 cm) coated with PTFE to minimize surface absorption of water and NH3. The SH contained the PS1, TB600B and MELBA sensors. An external pump drew air through at 5.5 l min-1 with a 30 cm PVC inlet allowing co-location of the air sampling (residence time ~52 s). From 19/8/24, 10 l min-1 flow was used to assess the influence of higher airflow (residence time ~ 29 s). The ECtox and Scentroid DR2000 unit were placed next to the SH, directly sampling from ambient air. The ECtox was installed so its inlet probe initially oriented parallel to the ground, however on the 19/08/2024 it was rotated 45° towards the release system to enhance the exposure to the air stream released from the high NH3 emission system. The DR2000 unit was placed under a weatherproof box with a PFA inlet <10 cm and a 0.8 l min-1 flow (residence time ~0.2 s).

Detailed information on the setup can be found in Espina-Martin et al. (2025).

# Collection/Generation/Transformation Methods

# Nature and Units of recorded values

### Monitoring site

The data was collected at Whim Bog, an ombrotrophic blanket bog located 11 km south of Edinburgh (NT 204532) 3◦16’ W and 55◦ 46’ N. Whim Bog has been operated as an experimental nitrogen enhancement site since 2002, where ecosystem and species responses to N pollution (wet and dry deposited) are investigated. The site includes a dry NH3 gas enhancement transect with an ammonia plume dispersing over a distance of ~100m (Leith et al., 2004; Sheppard et al., 2004 & 2011).

### NH3 sensors

One high resolution NH3 analyzer and six NH3 miniaturized sensors were compared.

- The Picarro G2013 NH3 Cavity ring down spectroscopy (CRDS) analyser uses infrared absorption. A 20 cm PFA inlet with a PTFE guard filter (0.001 µm porosity) sampled air at 1.5 l min-1 (residence time ~ 0.22 s). The instrument NH3 response (linearity and intercept) was checked in the laboratory before and after the deployment. A 100 ppm NH3 cylinder (N2 99,99 %, NH3 0.01%; BOC) was diluted with N2 (99.998% purity, BOC) through mass flow controllers (MFCs) (Bronkhorst Ltd, Netherlands). The Picarro prior and after deployment demonstrated excellent linearity (R2=0.999) across the whole concentration range. The slope of the theoretical to measured concentrations was 0.90 (pre-deployment) and 0.87 (post-deployment). Prior to the deployment on Whim Bog, new guard PTFE filters (0.015 µm pore size, Entegris) were placed to protect the Picarro analytical cavity. They were kept through the pre-campaign (no use) to the post-campaign (three months of use) linear check to assess if the exposure to high concentrations of NH3 during the intercomparison impacted the Picarro’s response. Picarro NH3 concentrations were not corrected to the theoretical due to potential losses of NH3 in the set-up used to check the linearity of the instrument.

- Six miniaturized sensors were tested in this experiment: TB600B-NH3-10, PS1-10-NH3 ECtox-50, AM1 LC, AM2 HC, and MELBA.

The PS1 and TB600B are 3-way electrode EC sensors attached to electronic boards that enable USB connectivity with a laptop running the supplier software for data recording and extraction. An external pump is required to pass air over the sensitive surface. The TB600B is a “ready to be used as soon as is connected” sensor. The PS1 is sold as such, but it requires a two-point calibration before use (zero and span) to be carried out by the user. This was done in a laboratory chamber (0.28 l, polycarbonate) with N2 (99.998% purity, BOC) for the zero and a 100 ppm NH3 (N2 99,99 %, NH3 0.01%; BOC) diluted with N2 used for the zero to achieve a span of 1 ppm. The supplier does calibrate the sensor in humid conditions; however, this was not possible to achieve due to laboratory set-up limitations. The PS1 and TB600B sensors have an adjustable reporting time from one second up to one minute. The sampling period for the campaign was set to 1 minute.

The ECtox sensor consists of a 16 cm-stainless steel probe with one end covered by a PTFE membrane that protects the sensing chamber. It samples air at 2 ml min-1 and analyses for one minute. The ECtox sensor used the calibration version of the TB600B software to adjust the sampling time and record data.

The AM1 and AM2 sensors come as part of the Scentroid DR2000R air lab product, a “black box” device equipped with T, RH, and seven atmospheric pollutant sensors designed to be installed into unmanned aerial vehicles for urban aerial air quality surveys. It samples air at 0.8 l min-1 from the inlet located in front of the unit. Data collected by the DR2000R box is transmitted to a tablet-based ground station through wireless connection using long range protocol in real-time. The ground station is equipped with visualization and processing software to visualize and export the data, however if the DR2000R connection is lost, data recording stops and needs to be restarted manually. The system battery and the ground station have a two-hour battery, requiring a continuous power supply to enable extended monitoring.

# Quality control

UKCEH staff members were visiting the site on a weekly basis to ensure the operation of the sensors and maximize data capture.

The Picarro is used as the “reference” with the assumption of a minimum of ±15% uncertainty, however previous tests18 show that the uncertainty lies within 5-7%. The Picarro data QA/QC involved checking the daily data streams and removing any data that would be out of the normal operational ranges of temperature and pressure within the internal cavity.

| Parameter | Valid Range |
| --- | --- |
| Internal cavity temperature | 130-150 Torr |
| Internal cavity pressure | 43-47 °C |

The sensors data was QA/QC ‘ed by removing the first 5 minutes of operation whenever they were restarted and removing any outliers coming from sampling malfunctioning.

# Details of data structure

| column | title | details |
| --- | --- | --- |
| 1 | Date_time | Date and time of the data collection in British Standard Time (BST) |
| 2 | Whim_Bog_active | Indicates whether the NH3 release system was on (1) or off (0) |
| 3 | Whim_Bog_windspeed (ms-1) | Wind speed measured at Whim Bog. Unit = m s-1 |
| 4 | Whim_Bog_Wind_direction (°) | Wind direction measured at Whim Bog. Unit = degrees (°) |
| 5 | Picarro_NH3_ppm | Atmospheric ammonia (NH3) gas concentration measured by the Picarro G2103; unit = ppm NH3 |
| 6 | Picarro_H20_(%) | H20 (g) vapour measured by the Picarro G2103; unit = % |
| 7 | TB600B_NH3_ppm | Atmospheric ammonia (NH3) gas concentration measured by the TB600B-10-NH3; unit = ppm NH3 |
| 8 | TB600B_T_TB600_(°C) | Temperature measured by the TB600B-10-NH3 sensor; unit = °C. |
| 9 | TB600B_RH_TB600_(%) | Atmospheric ammonia (NH3) gas concentration, unit = % |
| 10 | PS1_NH3_ppm | Atmospheric ammonia (NH3) gas concentration measured by the PS1-10-NH3 ; unit = ppm NH3 |
| 11 | PS1_T_SGX_(°C) | Temperature measured by the PS1-NH3-10 sensor; unit = °C. |
| 12 | MELBA_NH3_calc_ppm | Atmospheric ammonia (NH3) gas concentration measured by the MELBA sensor ; unit = ppm NH3 |
| 13 | Ectox_NH3_ppm | Atmospheric ammonia (NH3) gas concentration measured by the ECtox-50; unit = ppm NH3 |
| 14 | Ectox_t (°C) | Temperature measured by the SGX PS1-NH3-10 sensor; unit = °C. |
| 15 | Ectox_RH_(%) | Relative humidity measured by the ECtox-NH3-10 sensor; unit = % |
| 16 | AM2_HC1_NH3_ppm | NH3 measured by the AM2 high concentration NH3 sensor; unit = ppm NH3 |
| 17 | AM2_HC2_NH3_ppm | NH3 measured by the AM2 high concentration NH3 sensor; unit = ppm NH3 |
| 18 | AM1_LC1_NH3_ppm | NH3 measured by the AM1 low concentration 1 NH3 sensor ; unit = ppm NH3 |
| 19 | AM1_LC2_NH3_ppm | NH3 measured by the AM1 low concentration 2 NH3 sensor; unit = ppm NH3 |

## References

Leith, I., Sheppard, L., Fowler, D., Cape, J., Jones, M., Crossley, A., Hargreaves, K., Tang, S., Theobald, M., Sutton, M., 2004. Quantifying Dry NH3 Deposition to an Ombrotrophic Bog from an Automated NH3 Field Release System. Water, Air, & Soil Pollution: Focus 4, 207–218. https://doi.org/10.1007/s11267-004-3031-3

Sheppard, L.J., Crossley, A., Leith, I.D., Hargreaves, K.J., Carfrae, J.A., van Dijk, N., Cape, J.N., Sleep, D., Fowler, D., Raven, J.A., 2004. An Automated Wet Deposition System to Compare the Effects of Reduced and Oxidised N on Ombrotrophic Bog Species: Practical Considerations. Water Air Soil Pollut: Focus 4, 197–205. https://doi.org/10.1007/s11267-004-3030-4

- Sheppard, L.J., Leith, I.D., Mizunuma, T., Neil Cape, J., Crossley, A., Leeson, S., Sutton, M.A., van Dijk, N., Fowler, D., 2011. Dry deposition of ammonia gas drives species change faster than wet deposition of ammonium ions: evidence from a long-term field manipulation. Global Change Biology 17, 3589–3607. https://doi.org/10.1111/j.1365-2486.2011.02478.x

- Espina-Martin, P., Leeson, S.R., Nicoll, R., Pearson, C., Hernandez, C.M., Redon, N., Mullinger, N.J., Yeung, K., Twigg, M.M., Deshpande, A.G., Jones, M.R., Costello, H., Spelman, G., Braban, C.F., 2025. Are Ammonia Sensors Ready for Outdoor Use? ACS Omega. https://doi.org/10.1021/acsomega.5c07873